10MHz frequency source

Possible sources

All sources exist in ResLab.

Passive Hydrogen Maser

OCXO

OXCO 5 MHz

OXCO 10 MHz (broken)

Rubidium Frequency Standard

GPS receiver

Marconi Synthesizer, Model: 2024

Selfmade distributor [Klaus]

K+K distributor (not for high quality distribution)

HighTec Fiber ???

CSAC Chip für stabile 10MHz Referenz

Current Setup [30.05.2017]

Current 10MHz setup.pdf

Future Setup

Raumpläne IQO

10_MHz_Verlegung_IQ_08_2017

Experiment Ansprechperson Experiment Ansprechperson Kabelverlegeung	Raumnummer	Number of cables [new = , old = ] Mean cable Lab cable	Number of distributors (Number of outputs)	Cable length + “Lab cable Length” [m]	Holes [Number, size, place]	Notes
ResLab S. Sauer S. Sauer	D215	2 [] 16 []	2 (16)	1 (OXCO-Mean distributor) + 1 (Mean distributor - Reslab distributor) + 16*5	1, Ø=?, Reslab-Q-Port	-
Atom-Lab-Mg S. Sauer W. Friesen	D320	2 [] 8 []	1 (16)	2*[ 25(Reslab - Mg-Lab) + 10(Mg-Lab)] + 8*5	1, Ø=5-10cm, Q-Port-Atom-Lab-Mg	-
KRb J. Peise A. Igel	E311	1 + ? []	1 (16)	25(Reslab - Mg-Lab) + 20(Mg-Lab) + 60(Mg-Lab-KRB)	-	* KRB_10MHz_Kabel_08_2017 * Die Durchführungen sind alle groß genug, nur auf höhe des Elektroniklabors liegen die Kabel in der Nähe von Heizungsrohren. * Im Flur befindet sich ein offener Kabelkanal.
QAI,EPR J. Peise C. Pür	D313 → D315	1 + ? []	1 (16)	25 (Reslab - Mg-Lab) + 20(Mg-Lab) + 25(Mg-Lab-QAI) + 25(QAI-EPR)	-	-
ATLAS H. Albers H. Albers	D330	1 + ? []	1 (16)	25(Reslab - Mg-Lab) + 25(Mg-Lab) + 50 (Mg-Lab-ATLAS) + 5(ATLAS-Lab)	?, Ø=10cm, ?	* ATLAS_10MHz_Habel_08_2017 * Gibt keine Information wie es über den Seminarraum verlegt ist → Neue Löcher bohren oder zusätzlicher Durchgang zu VLBAI
VLBAI E. Wodey ?	D323	1 []	1 (16)	25(Reslab - Mg-Lab) + 25(Mg-Lab) + 20 (Mg-Lab-VLBAI) + 5 (VLBAI)	-	-
Quantengravi J. Mathias ?	D323	1 []	1 (16)	25(Reslab - Mg-Lab) + 25(Mg-Lab) + 20 (Mg-Lab-VLBAI) + 5 (VLBAI)	-	-
AG S.& C. Ospelkaus Mirko Sierke S. Grondwoski	E304A → E104A	1 [] 2 []	2 (min. 8)	25(Reslab - Mg-Lab) + 25 (Mg-Lab) + 100 (Mg-Lab - E-Trakt = vorhanden!) + E-Trakt	-	*AG_Ospelkaus_10MHz_Habel_08_2017 * Das vorhandene Kabel ist 100 Meter lang und hat bei 10 MHz eine Dämpfung von etwa 4,3 dB. * Kabelnanfang im Mg-Labor ist mit “CW Tiemann” und das Ende im E-Trakt ist mit “MG ERT” beschriftet
Q-Port S. Bode ?	1 []	-	-	-	Vorhandene Kabel ist verlegt, muss noch getestet werden
Lasus Thejs ?	D220	1 []	-	-	-	* Vorhandene Kabel ist verlegt, Aircell 7, muss noch getestet werden * Sein vorhandener Verteiler ist ein SRS FS710
Maius D. Becker	keine Antwort	keine Antwort	-	-	-
AG Morgner B. Kreipe	0	0	-	-	-
Gesamt	7 (16) 2 (min 8)	622

To Do

1)

“GPS →” Rb-10Mhz [right Output] → BNC cable → BNC-I-connector → BNC-SMA-connecotor → long cable → PNA

Channel J1 and J2: OCXO → short cable → SMA-I-connector → long cable → PNA

J1 → short cable → SMA-BNC-connector → TP [6MHz] → long cable → PNA

J1 → short cable → SMA-BNC-connector → TP [6MHz] → Doubler [FD-2+] → BNC-SMA-connector → long cable → PNA

J1 → short cable → SMA-BNC-connector → TP [6MHz] → Doubler [FD-2+] → BNC-SMA-connector → BP [SBP-10.7+] → long cable → PNA

J1 → short cable → SMA-BNC-connector → TP [6MHz] → Doubler [FD-2+] → BNC-SMA-connector → BP [SBP-10.7+] → BNC-SMA-connector → Att. [10dB] → BNC-SMA-connector → long cable → PNA

Overall setup: J1 → short cable → SMA-BNC-connector → TP [6MHz] → Doubler [FD-2+] → BNC-SMA-connector → BP [SBP-10.7+] → BNC-SMA-connector → Att. [10dB] → short BNC-cable → Amplifier [ZFL-500LN-BNC+] → LP [BLP-15] → Att. [20dB] → Att. [3dB] → BNC-SMA-connector → long SMA cable → PNA

Currently not possible.

2)
3)
4)
5)

Distibution

Timetech distributor: Timetech distributor_datasheet

Spectra dynamics distributor: HPDA-15RMi Datasheet_HPDA-15RMi

Feature	TimeTech	Spectra dynamics HPDA-15RMi
Phase Noise [dBc]	@ 10MHz 1Hz -148 10Hz -160 100Hz -165 1000Hz -166 10000Hz -166	@ <hi #ed1c24>???</hi> MHz 1Hz -155 (max -150) 10Hz -165 (max-160) 1000Hz -170 (max 168) 10000Hz -171 (max -170)
Channel/Channel isolation [dB]	-	Out-out: 130 (min 130)\\Out-in: 140 (min 145)
Groundloops	<hi #ed1c24>???</hi>	<hi #ed1c24>???</hi>
RF-Bandwidth [MHz]	5 to 100 MHz sine wave useable from 100kHz to 200 MHz	+/- 1 dB 0.5-65 (min 1-50)
Input Level [dBm]	Input level working: +6 to +17 (+14 dBm @ 100 MHz) Input level nominal: +13	1 dB compression 18 (min 17) @ <hi #ed1c24>???</hi> MHz
Output Level [dBm]	<hi #ed1c24>???</hi>	+18 @ <hi #ed1c24>???</hi> MHz
Gain [dB]	Gain: 0 (+0.15/-0.1 dB, 1 – 100 MHz, +0.15 to –0.5 dB @ 200 MHz) Gain variation: < 0.1, output vs. output	@5MHz 0.2 (max 0.5)
Impedance [Ω]	Input: 50, nominal Output: 50, nominal	Input: 50 Output: 50
Return loss [dB]	Input: >45 @5MHz and 10 MHz, >30 @100 MHz Output: >50 @5MHz and 10 MHz, >30 (typ –35) @ 100 MHz	Input(S11): -35 (max -30) Output(S22): -35 (max -30)
Distortion [dBc]	<hi #ed1c24>???</hi>	+13 dBm: -48 (max -45) +17 dBm: -42 (max -40)
Temp. Delay Coefficient	+15…+30°C 1K pp with maximum slopes 0.5K/h	0-50°C 1.2 (max 1.5) [ps/ºC]
Connectors	only BNC <hi #ed1c24>???</hi>	SMA for input and output
Power connectors	Power Consumption (all outputs loaded): AC: < 0.2 A Priority on AC supply input DC: < 0.9 A, Stand-By: < 50 mA DC is used as back-up supply Supply voltage AC: 100 to 240 V AC, 50/60 Hz, LED indicating AC voltage Supply voltage DC: 18 to 36 V DC, DC isolated, LED indicating DC voltage 2 redundant DC power input connectors, automatic switchover	AC Operation: 100-240 VAC, 47-63Hz DC Operation: 12-36 VDC-Available for some models
Number of Outputs	16 or 15	5-15
Number of Inputs	1 or 2	1-3
More	-

6)
7)

Which cables?

Best choise (cost-effectiveness) < 150 MHz: RG223 - DatasheetRG223 - conrad

Aircell 5 for > 100 MHz, low attenuation for higher frequecies: Aircell 5 - Datasheet Aircell 5 - conrad

For high frequencies <10 GHz: accuphasebrochure.pdf

Testing cable:

Offenes Ende vermessen, dann sollte es Hochohmig sein

Auf einer Seite mit 50 Ohm verbinden und vermessen: 50 Ohm+Kabelwiderstand (<3-4 Ohm ist ok)

Wenn der Kabelwiderstand höher ist, dann wurden vermutlich beim Crimpen (Luft-)Übergangswiderstände erzeugt

Which connectors?

BNC or SMA?

8)

OCXO

Vibration isolation for the OXCO?

Frequency double, how?

Filters

5MHz

10MHz