Checklist after power failure

Check both IGP Controller for the resonators R1 + R2

Current should be below 2 mA (both should show “HV ON” (green LED on)

Check both USV for IGP´S

Turn on the computers

Check Meerstetter TEC for R1 + R2 + ECDL

Green LED on?

Check in the software TEC Service v2.70 - R1, TEC Service v2.70 - R2 and TEC Service v2.70 - ECDL Baseplate on the ResLab-PC:

How is the behavior of the temperature?

Check both USV for TEC´s

Check the air conditioner for the entire room

Is the air-conditioning system switched on?

Is the V-belt (=Keilriemen) broken?

What is the temperature of the draft (=Zugluft)? → should be >20 °C

Check the electric fuse (=Sicherung) on the third floor

Check the H-Maser

All LED are green?

Check in the Software VCH on the ResLabMobil:

All values except for D2h must be green!

Check the 914nm-laser

Light is on? Powermodule is on?

TA stabilization is ok?

Laserdiode is on?

How much power is in front of the resonator? → should be 70-100 µW

What wavelength does the Wavemeter display?

Check the 10 MHz reference and set all frequencies in the lab

GPS-Receiver is working?

LED flashes orange/yellow?

If not Turn on the PWR switch until LED flashes

Start the software Net View and connect the receiver (COM Port2)

Rb-clock is working?

Check the LED´s:

Power-LED and Locked-LED have to be green

1 pps sync-LED has to be green

1 pps input-LED has to flash

Check the 10 MHz-frequency output with an spectrum analyzer, if there is 10 MHz.

Check the 10 MHz distributor synthesizer

Ext refernce 10 MHz direct is working?

Set correct values

Check all other synthesizer

Ext refernce 10 MHz direct is working?

Set correct values

Check frequency comb

Laser and amplifier are on?

Check LED´s

Do the software fiber laser and fiber comb control work?

Can you speak with the comb?

Check the counter

Does the software FXQE80 work?

If not: Turn off and on the counter (cold-device plug disconnect and connect)

Check the fiber Link to PTB

Turn everthing on, like ther spectrum analysizer

Restart the PC in the Atom Lab

Call them!

If the temperature is normal again, then…

Stabilize the laser of R1

Turn on the Laserdiode

Turn on the TA

Couple the light into the TA

Turn on the TA stabilization

Lock R1

Stablize the frequency comb

Lock Ceo

Lock Rep

Measure the frequencies with the counter:

914nm

frep

10 MHz

virtuel beat