# Synthesizer
## Important Synthesizers for ResLab
Typ Synthesizer	Application	Frequency [Hz]	RF-Power [dBm]	Reference	Comment
Marconi 2024	Reprate	1 000 000 226	7	Passive H-Maser
Marconi 2024	Amplifier for 10 MHz distribution	10 000 000	14	Rb Clock
Marconi 2024	Stell-AOM Mixdown	74 940 000	7	Rb Clock
Marconi 2024	Free				was broken
Marconi 2023B	Free
Rhode&Schwarz SMC100A	Free
Rhode&Schwarz SMC100A	30m Fiber stabilization Reference	160 222 000	4	Rb-Clock
Rhode&Schwarz Hameg HM8134-3	30m Fiber stabilization	80 111 000	1.4	Rb-Clock	* ext. Mod. on (400 kHz Dev.) * Name: AOM2 * Shit stability
Rhode&Schwarz Hameg HM8134-3	Free				Shit stability!
MG520 M Frequency Generator	EOM R1	13 810 000	10	Rb-Clock
Rhode&Schwarz SMH 845.4002.52	EOM R2	16 200 000	13	Rb-Clock
Rhode&Schwarz SML01	Double-pass AOM R1	104 630 000 + Computer	13	Rb-Clock	* Name: AOM0 * located in upstairs-Lab
Rhode&Schwarz SMC100A	* Linewidth Lock with Blue AOM R1 * 5m Fiber stabilization	77 111 000	2	Rb-Clock	* Name: AOM1 * located in upstairs-Lab
Rhode&Schwarz SMY01	5m Fiber stabilization Reference	77 111 000	5	Rb-Clock	* Name: AOM1 * located in upstairs-Lab
	Double-pass AOM R2

## Specks
Typ Synthesizer	Phase Noise @ 1 GHz [dBc]	Resolution [Hz]
Marconi 2024	-90 @ 100 Hz -100 @ 1 kHz -115 @ 10 kHz -121 @ 20 kHz	0.01 < 100 Hz 0.1 < 1 kHz 1 < 20 kHz
Marconi 2023B	-75 @ 100 Hz -100 @ 1 kHz -112 @ 10 kHz -121 @ 20 kHz	1
Rhode&Schwarz SMC100A	-45 @ 1 Hz -70 @ 10 Hz -82 @ 100 Hz -100 @ 1 kHz -115 @ 10 kHz with internal OCXO	0.1
Rhode&Schwarz Hameg HM8134-3	-15 @ 1 Hz -50 @ 10 Hz -80 @ 100 Hz -100 @ 1 kHz < -100 @ 10 kHz	1
Rhode&Schwarz SML01	-95 @ 1 Hz -120 @ 10 Hz -130 @ 100 Hz -138 @ 1 kHz -148 @ 10 kHz with 9.5 MHz output frequency -128 @ 20 kHz with 1 GHz output frequency	0.1

## Marconi Synthesizer
### Marconi Synthesizer 2024

- Manual 2024: ifr_2024_datasheet.pdf
- possible settings: Frequenzbereich von 9 kHz bis 2,4 GHz, Auflösung 1 Hz, Ausgangspegel von -140 bis +13 dBm, Auflösung 0,1 dB, interne 2-Ton-Modulationsquelle für Sinus, Rechteck und Dreieck, lineare und logarithmische Wobblung, AM, FM, pM, Puls und FSK, Rückleistungsschutz bis 50 W, LC-Display, nummerische Tastatur, GPIB und RS-232
Options
01 No attenuator
03 high Power
04 high stability frequency standard
11 Fast pulse and high power
- Error message: “External reference is missing!!!”

- Important: External(Ext): 1 Ext 10MHz direct
(MENÜ)-Button → Frequency/Sweep (Select) → 4 Frequency standard (Select) → 1 Ext 10MHz direct
- Possible errors: no Rb-referenz/ GPS/ OXCO connection or wires connected?
- Normal state (10MHz):

### Synthesizer is broken?
Manual for repair: http://www.ko4bb.com/getsimple/index.php?id=manuals&dir=Marconi
Quick repair project (Marconi 2024): a forum where severel defects are presented and explained how to fix them: http://www.eevblog.com/forum/repair/quick-repair-project-marconi-2024/
Output signal (dBm) is too small: Probably the amplifier is broken. Check backwards the signal inside the synthesizer with a measuring tip.
1. Adjust the frequency of the half from the maximal Oszi-frequency
2. Measure the distance between the minimum and maximum sinus-signal. Half of the distance devide by root(2) provide the effective voltage
3. Use the data sheet for dBm calculation: dbm-volts-wattconversion.pdf
Solved Problems:
Problem: No Output Signal
Solution: Transitor TR708 was broken. The Transistor has been replaced by a BC847 Transitor.
### Marconi Synthesizer 2023B (ehemals ausgeliehen von Q-Port)

Manual 2023B:datenblattmarifr2023b_2025.pdf
## Rohde&Schwarz Synthesizer
### Rohde&Schwarz Synthesizer SMC100A

- Manual SMC100A:smc100a_operatingmanual_en_09.pdf
- Datasheet SMC100A:smc100a_dat-sw_en.pdf
### Rohde&Schwarz Synthesizer SML01

Datasheet SML01:rohde-schwarz-sml01-datasheet.pdf
### Rohde&Schwarz Synthesizer Hameg HM8134-3

Manual Hameg HM8134-3:hm8134-3_usermanual_de_en_04.pdf